## Slide 1
Magnetique resonance imagering
Created bay :
Iles Amel

## Slide 2: Overview of lecture
Lecture 1: MRI equipment
Lecture 2: Spin Physics
Lecture 3: Relaxation time phenomenon
Lecture 4: MR Image reconstruction
Lecture 5: Pulse sequences
Lecture 6: Image contrast
Lecture 7: Parallel imaging
Lecture 8: Functional imaging (FMRI)
Lecture 9: Diffusion MRI

## Slide 3: MRI equipment
Chapter 1

## Slide 4: Introduction
Brief overview of the history of MRI
In 1946, Felix Bloch and Edward Purcell independently discovered the phenomenon of nuclear magnetic resonance (NMR).
They were awarded the Nobel Prize in Physics in 1952 for their work, which laid the foundation for the development of MRI.
In the 1970s, Raymond Damadian, a physician and scientist, introduced the concept of using NMR to detect cancerous tissue in the human body. He developed the first MRI scanner, which he conducted the first full-body MRI scan in 1977. However, his machine had limitations and was not as efficient as modern MRI scanners.
Paul Lauterbur and Sir Peter Mansfield made groundbreaking contributions to MRI. In 1973, Lauterbur introduced the concept of using gradients in magnetic fields to spatially encode the NMR signals, enabling the production of detailed images. Mansfield further developed this idea and introduced the use of rapid imaging techniques, such as echo-planar imaging (EPI), which greatly reduced the imaging time.
In the 1980s, advancements in technology and imaging techniques led to significant improvements in MRI. The introduction of superconducting magnets and the development of pulse sequences, such as the spin-echo sequence, made MRI images clearer and faster.
1990, fMRI
In 1999, K. P. Pruessmann developped parallel MRI technique.

## Slide 5: MRI equipment
The MRI consists of different rooms:
the waiting room,
the patient preparation room.

## Slide 6: MRI equipment
The scan (examination) room :
the MRI machine is located in the examination room.
The control room : is equipped with computers,
Some computers are used for MRI image acquisition: Manipualtors use these computers to reconstruct images of different planes: sagittal, coronal, and axial.
Other computers are used by radiologists for image processing of the acquired images.

## Slide 7: MRI equipment
Components required to power the MRI machine :
Components powering RF and gradient coils,
Cooling system (système de refroidissement).
(a)
RF Power Amplifier: Amplifies the radiofrequency (RF) signal used to excite the protons in the patient's body for imaging.
Gradient Power Amplifiers: Provide power to the gradient coils, which generate magnetic field gradients used for spatial encoding in the MRI images.

## Slide 8: MRI equipment
(b)
(b)
Chillers: Maintain optimal temperature for various components of the MRI system, such as the gradient coils, RF amplifiers, and cryogenic systems.
Heat Exchangers (échangeur de chaleur): Dissipate heat generated by the MRI machine and its components, ensuring efficient operation and preventing overheating (évite la surchauffe).

## Slide 9: components of the MRI equipment

## Slide 10: Faraday cage
In the context of an MRI facility, a Faraday cage may be incorporated into the design of the scanning room or magnet room. It helps to prevent external electromagnetic signals, such as radiofrequency interference (RFI) or electromagnetic radiation, from affecting the sensitive MRI equipment and distorting the acquired images. The Faraday cage enhances the overall signal quality and reduces artifacts during the MRI procedure.

## Slide 11: Faraday cage
The scan room of a magnetic resonance imaging (MRI) machine is designed as a Faraday cage.
The Faraday cage in an MRI facility typically consists of conductive walls, floors, and ceilings (plafond), which are carefully constructed and grounded to ensure proper shielding and containment of electromagnetic fields. This helps maintain a controlled electromagnetic environment, allowing for accurate and reliable imaging.

## Slide 12: Quench

## Slide 13: Shimming
One of the requirements for good imaging is a homogeneous magnet field. This is a field in which there is a uniform field strength (force) over the image area. Shimming is the process of adjusting the magnetic field to make it more uniform.

## Slide 14: Shimming
Inhomogeneities are usually produced by magnetically susceptible materials located in the magnetic field. The presence of these materials produces distortions in the magnetic field that are in the form of inhomogeneities. This can occur in both the internal and external areas of the field. Each time a different patient is placed in the magnetic field, some inhomogeneities are produced. There are many things in the external field, such as building structures and equipment, that can produce inhomogeneities. The problem is that when the external field is distorted, these distortions are also transferred to the internal field where they interfere with the imaging process. Inhomogeneities produce a variety of problems.
It is not possible to eliminate all of the sources of inhomogeneities. Therefore, shimming must be used to reduce the inhomogeneities. This is done in several ways. When a magnet is manufactured and installed, some shimming might be done by placing metal shims in appropriate locations. Magnets also contain a set of shim coils. Shimming is produced by adjusting the electrical currents in these coils. General shimming is done by the engineers when a magnet is installed or serviced.

## Slide 15: Magnet and coils

## Slide 16: Magnet
Superconducting magnets use coils cooled (refroidir) to very low temperatures, typically close to absolute zero, to generate a powerful and stable magnetic field. The manufacturing and cooling of these magnets, as well as the required safety measures, significantly increase their cost.
The unit of measurement for the magnetic field in MRI is the Tesla (T). 1 Tesla (T) is equal to 10,000 Gauss.
Note: Gauss is a unit of magnetic induction commonly used in some contexts, but Tesla is the standard unit used in the field of MRI.
A small ferrite-type metal magnet typically has a remanent field on the order of 2000 to 4000 Gauss.

## Slide 17: Magnet
    1. Permanent magnet
A permanent magnet is composed of a ferromagnetic material structure.
Note: Ferromagnetic materials, such as iron, nickel, and cobalt, have the ability to retain a significant amount of magnetization even after the external magnetic field is removed. This property allows them to be used as permanent magnets, which can generate their own magnetic fields without the need for an external power source.
There are different types of magnets used in MRI, such as permanent magnets, resistive magnets, and superconducting magnets. Among them, superconducting magnets are the most commonly used due to their superior performance.

## Slide 18: 1. Permanent Magnet
Its main advantages are:
Absence of eddy currents: Permanent magnets do not generate eddy currents, which are induced currents that can cause energy losses and heating in conducting materials.
No energy consumption: Permanent magnets do not require an external power source to maintain their magnetic field. Once magnetized, they can sustain their magnetic properties indefinitely without the need for ongoing energy input.
Disadvantage :
is the inability to quickly suspend or turn off the magnetic field in case of an emergency: "quench" in MRI systems, where the magnetic field needs to be rapidly and safely deactivated. In permanent magnet systems, it is more challenging to rapidly dissipate the stored magnetic energy, making emergency shutdowns more complex.
they typically produce a relatively lower magnetic field strength compared to superconducting magnets. The magnetic field strength of permanent magnets is often in the range of 0.4 Tesla or lower, whereas superconducting magnets used in high-field MRI systems can generate much stronger magnetic fields, ranging from 1.5 Tesla to 7 Tesla or more.
permanent magnets can be relatively heavy due to the materials used and the need for a larger volume to generate a sufficient magnetic field.

## Slide 19: 2. Resistive Magnet
    A resistive magnet is an electromagnet composed of a copper coil through which an electric current flows, producing a magnetic field at its center.
Advantages:
  Resistive magnets are relatively inexpensive compared to superconducting magnets because they do not require liquid helium for cooling.
  In case of an emergency, the magnetic field can be stopped within seconds by simply stopping the current flowing through the coil.
Disadvantages:
  Resistive magnets consume significant electrical power to sustain the current in the coils, which can lead to higher energy consumption compared to other magnet types.
  They are highly sensitive to temperature variations, which can result in magnetic field inhomogeneity. Maintaining a stable temperature environment is crucial for achieving consistent and accurate magnetic field strengths.
  Resistive magnets typically produce relatively lower magnetic fields, with maximum field strengths reaching around 0.5 Tesla or less. This limited field strength can be a drawback in applications that require higher field strengths for improved imaging resolution and sensitivity.

## Slide 20: 3. Superconducting Magnet
The basic components of a superconducting magnet include:
Superconducting Wire: Superconducting magnets rely (reposent) on superconducting materials, typically in the form of wires or coils, to carry the electrical current without resistance. These wires are made of materials that exhibit (présentent) superconductivity at low temperatures, such as niobium-titanium (NbTi) or niobium-tin (Nb3Sn) alloys (les alliages).
In superconducting magnet systems, typically liquid helium (either helium-4 or helium-3) is used for cooling the superconducting wires. The specific temperature at which the superconducting wire is cooled depends on the type of superconducting material used, with niobium-titanium (NbTi) typically requiring a temperature around 4 Kelvin, and niobium-tin (Nb3Sn) requiring a lower temperature of around 2 Kelvin or less.

## Slide 21: Superconducting Magnet
Advantages:
High Magnetic Field Strength: Superconducting magnets can generate significantly higher magnetic field strengths compared to other magnet types. They can achieve field strengths ranging from 1.5 Tesla to 7 Tesla or even higher, enabling improved image resolution and sensitivity in MRI.
Energy Efficiency: Superconducting magnets operate in a state of near-zero electrical resistance, which means they can sustain the magnetic field without the need for continuous electrical power input once cooled to their superconducting temperature. This results in lower energy consumption and reduced operating costs compared to resistive magnets.
Stability and Homogeneity: Superconducting magnets offer excellent field stability and homogeneity, which are crucial for obtaining high-quality and consistent MRI images. They can maintain a stable magnetic field over long periods, allowing for reliable imaging and accurate quantitative measurements.
Quench Protection: Superconducting magnets have built-in quench protection systems. In case of an event called a quench, where the superconducting state is lost, these systems safely dissipate the stored magnetic energy, preventing damage to the magnet and ensuring the safety of the surrounding environment and personnel.

## Slide 22: Superconducting Magnet
Disadvantages:
Cost: Superconducting magnets are typically more expensive to manufacture and operate compared to other magnet types due to the complex and specialized technology involved, as well as the need for cryogenic cooling systems.
Cooling Requirements: Superconducting magnets require cryogenic cooling systems to maintain the superconducting temperature. This involves the use of liquid helium or alternative cryogens, which adds complexity and cost to the system's operation and maintenance.

## Slide 23: Gradient coils

## Slide 25
Gradient coils
Z Gradient Coil
variations in the field produced by gradients in direction z

## Slide 26
Z Gradient Coil

## Slide 27
In all MR scanners the z-gradient is produced using two coils carrying current in opposite directions as shown in the diagram. This is known as a Maxwell coil configuration. Maxwell coils produce an incremental field that is zero at magnet isocenter but increases linearly outward in both the +z and -z directions. When this is added to the constant (Bo) field, the result is a gradually increasing field along the z-axis.
Bobines de gradient
Bobines de gradient selon Z

## Slide 28: X gradient coil
The X-gradient coil produces a magnetic field gradient in the X direction, which allows for spatial encoding of the MRI signal along the left-right axis. This gradient is essential for creating detailed images with accurate spatial information.

## Slide 29: y gradient coil
Similarly, the Y-gradient coil generates a magnetic field gradient in the Y direction, enabling spatial encoding along the anterior-posterior axis. The Y-gradient coil contributes to the production of multi-planar images and helps determine the location of structures within the body from front to back.
Together, the X-gradient coil and Y-gradient coil work in conjunction with the Z-gradient coil to spatially encode the MRI signal and facilitate the formation of detailed, three-dimensional images of the imaged area.

## Slide 30: Radiofrequency (RF) coils
Certain coils can be transmitters, receivers or both.

## Slide 31: Radiofrequency (RF) coils
certain coils can be transmitters, receivers or both.
In transmission, the goal is to deliver uniform excitation throughout the scanned volume. On reception, the coils must be sensitive and have the best possible signal to noise ratio.
An MR scanner generally contains a « whole body » coil, located in the cylinder of the machine, uniformly covering the entire scan volume. The sensitive volume of surface coils, being placed in direct contact with the zone of interest, has less depth and is more heterogeneous. However, surface coils offer a better signal to noise ratio and imaging capacity with higher spatial resolution. The homogeneity and sensitive volume of surface coils can be improved by combining them into a phased array. They still have the advantage of a better signal to noise ratio, but at the cost of more complex signal processing.

## Slide 32: Radiofrequency (RF) coils
Volume coils: are transmitters and receivers coils, they cover all objet to be imaged. They emit the B1 field and receive the RF signal.

## Slide 33: Radiofrequency (RF) coils

## Slide 34: Radiofrequency coils (RF)
Surface Coil:
Increased Sensitivity: The surface coil is designed to be placed on or near the body surface directly over the region of interest. It enhances the sensitivity of the MRI signal from that specific area, improving the signal-to-noise ratio (SNR) and allowing for better image quality and detail.
High-Resolution Imaging: By focusing on a specific region, such as the head, spine, or extremities, the surface coil provides high-resolution imaging of that area. It captures fine anatomical details and can be particularly useful for imaging small structures or areas close to the surface.
Localized Imaging: Surface coils enable localized imaging, as they capture signals primarily from the region directly beneath or near the coil. This localized approach allows for more targeted imaging and better visualization of specific structures or pathologies in that area.

## Slide 35: Bobines radiofréquence (RF)
Volume Coil:
Whole-Body or Large Region Imaging: Volume coils are larger and cover a wider area, making them suitable for imaging larger regions or the entire body. They provide a homogeneous magnetic field and allow for comprehensive imaging of the body, including organs, tissues, and anatomical structures across a larger field of view.
Transmitting RF Pulses: Volume coils are often used as transmit coils, delivering the radiofrequency (RF) pulses needed to excite the protons in the body and generate the MRI signal. These coils emit a uniform RF field that covers the entire imaging volume, enabling efficient excitation of the tissue.
Receiving the MRI Signal: Volume coils also serve as receive coils, capturing the MRI signal emitted by the excited protons. They collect signals from the entire imaging volume, providing a comprehensive view of the region being imaged.
Balanced SNR: Volume coils offer good overall signal-to-noise ratio (SNR) throughout the imaging volume. While they may not provide the same level of sensitivity as surface coils in specific localized areas, they offer a balanced SNR across the larger region, allowing for reliable (fiable) imaging and diagnostic evaluation.

## Slide 36: Phased array coils
Phased array coils consist of multiple smaller coil elements arranged in an array or grid pattern. Each coil element can be individually controlled, allowing for dynamic adjustment of the RF signal's phase and amplitude. This phased array configuration offers several advantages:
Increased Signal Sensitivity: Phased array coils utilize multiple coil elements to capture the MRI signal simultaneously from different locations. This arrangement enhances signal sensitivity and improves the signal-to-noise ratio (SNR) by combining the signals received from each element.

## Slide 37: Phased array coils
Parallel Imaging and Accelerated Imaging: The ability to control individual coil elements enables parallel imaging techniques. By acquiring data from multiple coil elements in parallel, phased array coils can accelerate imaging acquisition time, reducing the overall scan duration. This is particularly useful for acquiring images with shorter breath-holds (respiration) or reducing motion (mouvement) artifacts.
Flexible Field of View and Spatial Resolution: Phased array coils allow for flexible field-of-view (FOV) coverage. The individual coil elements can be activated or combined to adapt the FOV and achieve higher spatial resolution in specific regions of interest.
Improved Image Quality and Uniformity: The combined signals from the phased array coil elements can be processed to create high-quality images with improved uniformity and reduced artifacts.

## Slide 39
In MRI, the system of detection refers to the components responsible for receiving and detecting the radiofrequency (RF) signals emitted by the excited protons in the body. The main components of the detection system in MRI include:
RF Receiver Coils: RF receiver coils, also known as RF receive coils or simply receive coils, are responsible for capturing and detecting the RF signals emitted by the excited protons. These coils are placed strategically around the region of interest to maximize signal reception. Depending on the imaging needs, different types of receiver coils can be used, such as surface coils, volume coils, or phased array coils.
Pre-Amplifiers: The pre-amplifiers are electronic devices that amplify the weak RF signals received by the coils. They boost the signal strength before further processing and digitization.
Analog-to-Digital Converters (ADC): ADCs convert the analog RF signals into digital format, making them suitable for digital signal processing and analysis.
Signal Processing Unit: The signal processing unit performs various digital signal processing tasks on the digitized RF signals. This includes filtering, noise reduction, Fourier transformation, and image reconstruction algorithms to convert the raw data into meaningful images.
Image Reconstruction Algorithms: These algorithms process the acquired RF signals and reconstruct the final MRI images. Different reconstruction techniques, such as Fourier Transform-based methods (e.g., Fourier Transform Imaging) or more advanced techniques like parallel imaging or compressed sensing, can be employed to generate high-quality images.