## Slide 1
Magnetique resonance imagering
Created by:
Iles Amel
بسم الله الرحمن الرحيم

## Slide 2: Chapter 3
Relaxation time phenomenon

## Slide 3: introduction
In 1946 Felix Bloch published a mathematical analysis of the "Nuclear Induction" phenomenon including a set of equations explaining the origin and properties of the NMR signal.
Bloch began with the assumption that the millions of individual nuclei in a sample could be represented by a single vector M: it is more frequently known as the net magnetization.
Since M is a vector rotating in space it can be resolved into have three components — Mx(t), My(t), and Mz(t) — each a function of time. Mx(t) and My(t) are the transverse components; Mz(t) is the longitudinal component.
For the case of simple precession of M in a cone around B, the longitudinal component is a constant, Mz(t) = C. The transverse components rotate sinusoidally around the z-direction at angular frequency ω.
ω = γ B

## Slide 4: Excitation and detection
Effect of external radiofrequency (RF) field B1(t)
Once magnetization is rotated into the transverse (x, y) plane : RF is removed
Precession is again
  • about B0
  • at the Larmor frequency (in stationary frame of reference)
  • according to Bloch equation
Can detect the rotating magnetization

## Slide 5: Excitation and detection
Bloch reasoned that NMR signal decay occurs because the individual spins that comprise M interact with each other and their environment. As such, M would not precess indefinitely around B, but would instead seek to return to its initial alignment parallel to B with equilibrium magnitude Mo. To do this, the spin system had to release its energy to the environment, a process Bloch termed relaxation.
Bloch introduced two relaxation time constants, which he called T1 and T2, to account for the reestablishment of thermal equilibrium of the nuclear magnetization after generation of the NMR signal. T1 reflected the regrowth of longitudinal magnetization (Mz), whereas T2 characterized the decay of the transverse components (Mx and My).

## Slide 6: Bloch Equations
During return to equilibrium, the equations of motion for M in simplified form after a 90°-pulse become
Mx(t) = Mo e −t/T2 sin ωt
My(t) = Mo e−t/T2 cos ωt
Mz(t) = Mo (1 − e−t/T1)
These equations predict that M will exhibit a spiraling precession around B at the Larmor frequency with decay of transverse components back to zero and regrowth of the longitudinal component to its original maximum value Mo. Because T2 is always shorter than T1, the transverse components typically decay completely before the longitudinal magnetization is fully restored. The spiral pattern is sometimes said to resemble a "tent".

## Slide 7: What is T1 relaxation?
The longitudinal relaxation time, T1, therefore represents the time required for Mz to grow from 0 to (1 − 1/e), or about 63% of its final value.
T1 relaxation is the process by which the net magnetization (M) grows/returns to its initial maximum value (Mo) parallel to Bo. Synonyms for T1 relaxation include longitudinal relaxation, thermal relaxation and spin-lattice relaxation. The meanings and implications of these synonyms will become apparent shortly (peu).
M reaches 63% of its maximum value (Mo) at t = T1 and is very close to maximal at t = 5 x T1.

## Slide 8: What is T2 relaxation?
Similarly, the transverse relaxation time, T2, represents the time required for Mx or My to decay down to 1/e (about 37%) of their initial maximum values. T2 relaxation is the process by which the transverse components of magnetization (Mxy) decay or dephase. Synonyms for T2 relaxation are transverse relaxation and spin-spin relaxation.

## Slide 9: T1 and T2 relaxations

## Slide 10: RELAXATION AND TISSUE CONTRAST
Typical T1 and T2 values in human tissues at 1.5 T
T1 and T2 differ between tissue types
T1 and T2 are field-strength dependent

## Slide 11: Why is T1 longer than T2?
1) that for most organs T1 values are typically about 5-10x longer than T2 values;
2) that pure liquids (water/CSF) have very long T1 and T2 values;
3) that dense solids (ice, tendons, and proteins) have very short T2 values.
But the most important determinant is the size and motion of the molecule on which the hydrogen nucleus resides. This can be appreciated in the graph right showing T1 and T2 relaxation times plotted as a function of molecular "tumbling" rate. Small, rapidly rotating molecules (like free water) have long T1 and T2 times. As molecular motion slows (as in proteins and dense solids), T2 shortens and T1 again increases.

## Slide 12
Where does the MR signal come from?
The MR signal is a small electrical current induced in the receiver coil by the precession of the net magnetization (M) during resonance. This is a manifestation of Faraday's Law of Induction, wherein a changing magnetic field induces a voltage in a nearby conductor.

## Slide 13
Where does the MR signal come from?
After the B1 field is turned off, M continues to precess around Bo. A receiver coil sensitive to magnetic flux changes in the transverse plan can then detect the Mxy components as they oscillate at the Larmor frequency. A small current is generated in this coil via the Faraday-Lenz induction principle.
The signal is a sine wave oscillating at the Larmor frequency (ωo).
The recorded signal from such a pulse experiment has the form of either a free induction decay (FID) or echo.

## Slide 14: NMR signal
The NMR signal, however, does not persist forever. The initially coherent transverse components of M dephase as a result of both magnetic field inhomogeneities and intrinsic T2 mechanisms, incorporated in the concept of T2*-decay.
The resulting signal is the FID, a damped sine wave of the following form
[sin ωot ] e-t/T2*
In our simple examples we tacitly considered the MR signals to have arisen from a single large voxel in a homogeneous sample without spatial encoding.

## Slide 15: NMR signal
The FID is just one of four basic types of NMR signals produced in different ways. These are listed below

## Slide 16: NMR signal
the MR signal is an induced current generated by precession of the net magnetization (M) after stimulation by an RF- pulse. The signal is commonly detected in quadrature using receiver coils sensitive to magnetic flux in two orthogonal directions. Output channels, denoted I and Q (for "in phase" and "quadrature" respectively), send their respective signals along separate digitization and amplifier pathways. These signals are ultimately demodulated, processed, and recombined to create the final MR image.

## Slide 17: NMR signal
The quadrature receiver coils are measuring the same precessing magnetization (M) from two different perspectives. The signals in the I and Q channels, therefore, should theoretically be identical except for a 90º-phase shift between them. The second coil permits knowledge of the exact position of M and hence its direction of its rotation (i.e., positive vs negative frequency). Another way to think of this is that with quadrature detection you are "listening" to the MR signal in "stereo".
In reality the signals from the I and Q channels are not phase-shifted exact copies of one another because they also contain noise. Unlike the signals, noise in the two channels is independent and uncorrelated. Thus quadrature detection offers an increase in signal-to-noise by a factor of √2 = 1.41 over detection by a single linear receiver coils.

## Slide 18: NMR signal
The MR signal can be represented as a vector with real (Re) and imaginary (Im) components recorded from the I and Q channels respectively. An equivalent/alternative representation of the signal is as a complex number
Signal = (Re, Im) = Re + i Im
The signal data may be reconstructed in several ways:
as a "real" image,
as an "imaginary" image,
as a magnitude image, or
as a phase image.

## Slide 19
For illustrative purposes, the data was processed from the I and Q channels separately to demonstrate the contribution of each component more concretely. The horizontal bands are phase-shifted from each other by 90° and reflect a baseline phase error across the imaged volume.
Remark: In clinical practice, separate I and Q channel images are never obtained, with magnitude images being used nearly exclusively for diagnosis. Phase-images are occasionally generated in clinical MRI for the depiction of flow and characterization of susceptibility-induced distortions.

## Slide 21: 2D Fourier transforms
The MR imaging process consists of two major functions:
The first is the acquisition of RF signals from the patient’s body and the second is the mathematical reconstruction of an image from the acquired signals.
During the acquisition process the signals are collected, digitized, and stored in computer memory in a configuration known as k space. The k space is divided into lines of data that are filled one at a time. One of the general requirements is that the k space must be completely filled before the image reconstruction can be completed.

## Slide 22: 2D Fourier transforms
In 2DFT imaging, each row in k-space corresponds to the echo data obtained from a single application of the phase-encoding gradient. By convention, rows near the center of the k-space grid are defined to correspond to low-order phase-encode steps, whereas those rows near the top and bottom correspond to higher-order phase-encodings. Since echo amplitudes are larger at the low-order phase-encode steps (there is less gradient-induced dephasing), the values of k-space will be greater near the center of the grid.

## Slide 23
The Fourier transform is a mathematical technique that allows an MR signal to be decomposed into a sum of sine waves of different frequencies, phases, and amplitudes.
Fourier showed that any periodic signal s(t) can be written as a sum of sine waves with various amplitudies, frequencies and phases
Joseph Fourier
where ai are amplitudes, ϕi are phase shifts, and ω is the fundamental frequency. The higher order frequencies 2ω, 3ω, etc. are called harmonics.
Fourier transform

## Slide 26
Fourier transformation is the mathematical procedure connecting s(t) and S(ω). If s(t) is specified, S(ω) may be computed, and vice versa. The equations require some knowledge of complex numbers and calculus to make sense:
The common use in RF-pulse design is the sinc function:
whose Fourier transform is a uniform band of frequencies, such as those defining a single slice in conventional 2D MR imaging.

## Slide 28: What are 2D Fourier transforms?
The same idea can be extended into 2D:
The diagram to the right shows how a 2D brain image could be decomposed into planar waves of different amplitudes and frequencies. The sum of these many waves is used to build up a planar image.
A 2-dimensional image can be decomposed into a set of planar waves varying in phase, frequency, amplitude, and direction.
Unlike linear waves in one dimension, 2D planar waves may occur in any direction across the image (left-to-right, top-to-bottom, at 30°, 129°, etc.) Thus the Fourier summation must include all these angles and directions to represent the full image.

## Slide 32: 2D Fourier transforms
Création du motif "Egg-crate" (Gaufrier)
Haut de la figure (Addition de deux ondes) : On voit l'addition de deux ondes planes (représentées comme des surfaces ondulées) :
  La première onde (bleue, gauche) ondule le long de l'axe horizontal.
  La deuxième onde (rose, droite) ondule le long de l'axe vertical.
  Chaque onde plane est une sinusoïde 2D (une composante de Fourier), caractérisée par sa fréquence et son orientation.
Bas de la figure (Résultat) : L'addition de ces deux ondes, décalées de $90^\circ$ en orientation, produit un motif en forme de "boîte à œufs" ou de "gaufrier"
("Egg-crate" pattern). Ce motif est plus complexe et contient plus de détails que les ondes individuelles.
Émergence de l'Image (Droite de la figure)
"As planar waves from more directions are added, the image begins to emerge."
Le cube de droite montre le résultat de l'addition de multiples ondes planes (plus de deux), chacune ayant sa propre fréquence et son orientation.
En ajoutant un nombre suffisant de ces ondes planes avec les bonnes amplitudes, le résultat est une structure localisée, ressemblant ici à un pic 3D (un $\text{sinc}$ ou un $\text{gaussien}$), entouré d'autres interférences.
Chaque point noir sur le plan représente l'intensité d'une onde plane spécifique dans l'espace de Fourier (le plan u, v). L'image finale (le pic) est la somme (superposition) de toutes ces ondes. C'est le principe de la synthèse de Fourier.

## Slide 33: K space
The cells of k-space are commonly displayed on rectangular grid with principal axes kx and ky. The kx and ky axes of k-space correspond to the horizontal (x-) and vertical (y-) axes of the image. The k-axes, however, represent spatial frequencies. Each k-space point contains spatial frequency and phase information about every pixel in the final image. Each pixel in the image maps to every point in k-space.

## Slide 34: K space

## Slide 35: K space
Points along the kx-axis represent frequency components along the x-direction of the image. Conversely, points along the ky-axis reflect frequency components along the y-direction of the image. The k-space image of a square block whose sides are aligned with the x- and y-axes would look like the top image (right).
If the block were rotated by 45º, the dominant spatial frequencies corresponding to the edges of the block would now be oriented at a 45º angle to the kx- and ky-axes. The relation between angle of the planar image waves and k-space location is shown below.

## Slide 36: K space
Real images will, have frequency components in all directions, resulting in a "galaxy"-like appearance of k-space. A bright "star" at a particular location in k-space implies that the spatial frequencies that point represents are making a positive contribution to final image.

## Slide 37: K space
The data to fill k-space is taken directly from the MR signal but can be acquired in any order. The dominant method for filling k-space over the last 30 years has been the line-by-line Cartesian method. Today spiral and radially oriented trajectories are becoming more popular.
In the Cartesian method each digitized echo completely fills a line of k-space. The echo signal is recorded in quadrature, so each k-space point contains real and imaginary components. The k-space values on the left side of each row are obtained early in the evolution of the echo, while those on the right side are obtained later. The center of the echo (and hence largest values) occur near the middle of each row of k-space.

## Slide 38: FOV and k-space