## Slide 1
Magnetique resonance imagering
Created by:
Iles Amel
بسم الله الرحمن الرحيم

## Slide 2: Chapter 5
Pulse sequences

## Slide 3: Introduction
There are several different imaging methods that can be used to create MR images. The principal difference among these methods is the sequence in which the RF pulses and gradients are applied during the acquisition process. Therefore, the different methods are often referred to as the different pulse sequences. An overview of the most common methods is shown in Figure. As we see, the different methods are organized in a hierarchy structure. For each imaging method there is a set of factors that must be adjusted by the user to produce specific image characteristics.
The selection of a specific imaging method and factor values is generally based on requirements for contrast sensitivity to a specific tissue characteristic (PD, T1, T2) and acquisition speed. However, other characteristics such as visual noise and the sensitivity to specific artifacts might vary from method to method.
All of the imaging methods belong to one or both of the two major families, spin echo or gradient echo. The difference between the two families of methods is the process that is used to create the echo event at the end of each imaging cycle.

## Slide 4: What are TR and TE?
The echo time (TE) represents the time from the center of the RF-pulse to the center of the echo. For pulse sequences with multiple echoes between each RF pulse, several echo times may be defined and are commonly noted TE1, TE2, TE3, etc. The repetition time (TR) is the length of time between corresponding consecutive points on a repeating series of pulses and echoes.

## Slide 5: The received signal depends on the longitudinal relaxation time T1 and the repetition time TR.

## Slide 6: Spin echo sequence
The simplest form of the spin-echo (SE) pulse sequence consists of 90°-pulse, a 180°-pulse, and then an echo.
The spin echo process showing the use of a 180° pulse to rephase the protons and to produce an echo event.

## Slide 7: Spin echo sequence
The time between the initial excitation and the echo signal is TE. This is controlled by adjusting the time interval between the 90˚ and the 180˚ pulses, which is 1/2 TE.

## Slide 8: Spin echo sequence
TR>> TE

## Slide 11
T=TE/2
180° pulse excitation
T=TE
rephasing

## Slide 12
The Spin Echo Process

## Slide 13: Inversion Recovery
Inversion recovery (IR) is a conventional spin echo (SE) sequence preceded by a 180° inverting pulse. In other words, if a SE sequence is denoted by {90°−180°−echo}, the IR sequence can be written:
180° — {90°−180°−echo}
The time between the 180° inverting pulse and the 90°-pulse is called the inversion time (TI).

## Slide 14: Inversion Recovery

## Slide 15: Tissue suppression
One application is to produce a high level of T1 contrast and a second application is to suppress the signals and resulting brightness of fat and fluids.
The signal is reset to zero by : TI ≈69% T 1

## Slide 16
Signal Cancellation in the Inversion-Recovery Sequence
The function of the inverting pulse is to flip the initial longitudinal magnetization (Mo) of all tissues in the imaged slice or volume to point opposite to the direction of the main magnetic field (Bo). During the TI interval, these inverted tissues undergo T1 relaxation as they variably seek to re-establish magnetization along the +z-direction. When spin echo signal generation begins (at the 90°-pulse), the initial longitudinal magnetizations of different tissues are now separated based on their different intrinsic T1 relaxation times. The degree of separation (and hence image contrast) is controlled by varying the TI parameter in the pulse sequence.

## Slide 17: T1-Based Fat And Fluid Suppression
Let us recall that fat has very short T1values (260 msec) and fluids have very long T1 values (2000 msec).
These values are outside of the range of the T1 values of other tissues in the body and are separate and not mixed in with the others. This makes it possible to use T1 as a characteristic for the selective suppression of both fat and fluid.
STIR Fat Suppression
STIR is an inversion recovery method with the TI adjusted to selectively suppress the signals from fat

## Slide 18: Flair sequence
The suppression of fluid by selecting a long TI that will image the longitudinal magnetization at the time when fluid is relaxing through zero

## Slide 19: STIR & FAT SAT
Suppress the Fat signals: fat STIR sequence

## Slide 21: Suppress the (LCR) in Flair sequence

## Slide 22: Gradient echo sequence
A gradient echo sequence is an MRI technique that uses a flip angle less than 90° for the RF (radiofrequency) pulse and does not require a 180° refocusing pulse. This sequence generates images by using magnetic field gradients to form echoes, rather than relying on spin refocusing. Gradient echo sequences are typically faster than spin echo sequences, making them suitable for dynamic imaging.

## Slide 27: Echo-Planar Imaging (EPI)
Although EPI may seem like a recent development, the technique is among the oldest methods of spatial localization in MRI, first described by Mansfield in 1977.
One of their early EPI techniques (MBEST) which became relatively popular in the 1980's is diagrammed below. Following a spin-preparation module (which might be as simple as a single RF-pulse), a strong switched frequency-encoding gradient was applied simultaneously with an intermittently "blipped" low-magnitude phase-encoding gradient. Gradient echoes (GREs) were collected with each oscillation of the readout (frequency) gradient.

## Slide 28: definition
Echo-Planar Imaging (EPI) is a fast MRI technique that allows for the rapid acquisition of images by capturing all the data needed for an image within a single excitation of the tissue. EPI is widely used in functional MRI (fMRI) and diffusion-weighted imaging due to its ability to capture images in a fraction of a second.
In a conventional MRI scan, multiple repetitions are required to fill k-space line-by-line, with each repetition corresponding to a new phase encoding step. This results in long acquisition times. EPI, however, speeds this up significantly by acquiring multiple k-space lines in quick succession, immediately after a single excitation pulse.

## Slide 29: Main Steps in EPI
RF Excitation Pulse The image acquisition begins with a radiofrequency (RF) pulse that excites the proton spins in tissues, aligning them in the transverse plane.
Rapid Echo Train Acquisition Right after the RF pulse, a series of magnetic field gradients are applied to capture "echoes" of the signal very quickly.

## Slide 30: Main Steps in EPI
the rapid series of magnetic gradients that are applied immediately after the RF pulse in Echo-Planar Imaging (EPI) mainly involve:
Frequency-encoding (Readout) Gradient: This is the primary gradient in EPI that oscillates rapidly back and forth during the echo train. Its role is to capture the signal echoes by switching directions, allowing multiple lines in k-space to be filled with each oscillation. This oscillation creates the sequence of echoes (echo train) that fills k-space along the frequency-encoding (or readout) axis in a zig-zag pattern.
Phase-encoding Gradient "Blips": Small adjustments in the phase-encoding gradient, called "blips," are applied between each echo. These blips shift the acquisition up or down to the next line in k-space, allowing the zig-zag filling pattern along the phase-encoding direction.
To summarize: the main gradient that oscillates is the frequency-encoding gradient, while the phase-encoding gradient applies smaller "blips" between echoes to move up or down in k-space.

## Slide 31: Acquisition en Série d'Échos dans l'Echo-Planar Imaging (EPI)
Dans l’EPI, une fois que l'impulsion RF a excité les protons dans le plan transverse, deux types de gradients sont utilisés pour capturer les lignes de données du k-espace très rapidement :
Gradient d'Encodage de Fréquence (ou Gradient de Lecture) Ce gradient, souvent appelé gradient de lecture, est le principal dans l’EPI. Il oscille rapidement d’un côté à l’autre, c’est-à-dire qu’il change de polarité plusieurs fois pendant une très courte durée. Cette oscillation permet de capturer une série d’échos successifs, créant ainsi le "train d'échos" (ou echo train). À chaque oscillation, une ligne entière dans le k-espace est remplie le long de l'axe de fréquence. Cet effet de va-et-vient permet d’acquérir plusieurs lignes du k-espace en très peu de temps.
Blips du Gradient d'Encodage de Phase Entre chaque écho, de petits ajustements appelés blips sont appliqués dans le gradient d’encodage de phase. Ces petits "blips" permettent de passer à la ligne suivante dans le k-espace. Ainsi, au lieu de capturer chaque ligne une par une avec une nouvelle impulsion RF (comme dans l'IRM conventionnelle), l’EPI utilise ces blips pour avancer dans le k-espace de manière continue en zig-zag.
Résumé des Rôles des Gradients
Le gradient de fréquence oscille pour capturer les échos en alternant de gauche à droite, remplissant les lignes du k-espace.
Le gradient de phase applique des blips entre chaque écho pour avancer dans le k-espace et capturer la ligne suivante.

## Slide 32: Echo-Planar Imaging (EPI)
Fast Filling of k-Space
The k-space is the raw data space of MRI. Unlike conventional MRI, which fills k-space line by line, EPI captures multiple k-space lines at once using a zig-zag pattern.
The first downward-going frequency-and phase-encoding lobes imparted initial negative phase shifts to all spins, moving the signal from the middle of k-space to the far left bottom corner to a point identified as (−kxmax,−kymax).
Les premiers lobes descendants d’encodage de fréquence et de phase appliquent des décalages de phase négatifs initiaux à tous les spins, déplaçant le signal du centre du k-espace vers l’angle inférieur gauche, en un point identifié comme (−kxmax,−kymax).

## Slide 33: Echo-Planar Imaging (EPI)
Subsequent positive and negative frequency-encoding lobes swept k-space from left-to-right and right-to-left respectively. Meanwhile, the blipped low-amplitude phase-encoding gradient pulses produced a step-wise increase along the ky-axis.
echo planar imaging referred to a sequence in which data from all of k-space for an entire 2D plane was collected following a single RF-excitation pulse. More recently the term has been expanded to include any rapid gradient-echo or spin-echo sequence in which k-space is traversed in one or a small number of excitations. In the modern lexicon these are termed single shot EPI and multi-shot EPI, respectively

## Slide 34: Single-Shot EPI vs. Multi-Shot EPI
Single-Shot EPI: The entire k-space is acquired in one go (or “shot”) following a single RF excitation pulse.
This approach is the fastest and enables capturing an image in milliseconds, but is more susceptible to artifacts
and distortions due to magnetic field inhomogeneities.
Multi-Shot EPI: Instead of capturing all k-space lines in a single shot, the acquisition is divided into multiple shots,
each acquiring a fraction of k-space. This improves image quality but reduces acquisition speed.
Some additional vendor-specific terminology accompanies these concepts. The number of k-space lines (echoes) collected in a single shot is called the "Echo Train Length (ETL)" by GE and Canon; the "EPI factor" by SIemens and Philips; and the "shot factor" by Hitachi. The regions of k-space divided up by each shot of a multi-shot sequence are called segments. Hence it is common to refer to lines (or views) per segment as a measure of the intensity of k-space coverage.
In a very broad sense, a fast (turbo) spin-echo sequence with a very high turbo factor/ETL could be considered a multi-shot EPI sequence. The HASTE/SS-FSE technique is an example of this.